PRODUCT REQUIREMENTS DOCUMENT (PRD)
RetinaSight — Explainable AI for Diabetic Retinopathy Screening

SIH 2026 • Problem Statement 26038 • MATLAB + Simulink

# 1. Product Overview

RetinaSight is a MATLAB-based screening and decision-support prototype for diabetic retinopathy (DR) in rural primary healthcare settings. It combines image quality assessment, adaptive enhancement, retinal/lesion analysis, DR severity grading, explainability, calibrated confidence, human review, and a Simulink model for district-scale deployment.

# 2. Product Goal

Prevent unreliable diagnosis from poor-quality fundus images.

Provide automated DR severity screening with clinically interpretable evidence.

Keep an ophthalmologist in the decision loop.

Demonstrate scalable telemedicine operations for 100,000+ patients/year.

Use MATLAB/Simulink capabilities meaningfully rather than as decorative components.

# 3. Target Users

PHC/field operator — captures images and follows recapture guidance.

Ophthalmologist — reviews AI output, evidence, and report; approves or overrides.

District program administrator — uses Simulink outputs for capacity and resource planning.

# 4. End-to-End Workflow

Fundus Camera → Image Quality Assessment → [Good / Borderline / Ungradable] → Borderline Enhancement + Quality Re-check → Retinal Structure/Lesion Analysis → DR Grading (0–4) → Explainability + Calibrated Confidence → Annotated Report → Ophthalmologist Review → Final Referral Decision.

Parallel deployment workflow: PHC → Image Acquisition → Network/Bandwidth → AI Processing → Queue → Ophthalmologist Review → Simulink throughput/capacity/resource analysis.

# 5. Functional Requirements

| ID | Requirement | Acceptance Direction |
| --- | --- | --- |
| FR-01 | Image ingestion | Accept standard fundus images (JPG/PNG) and initiate the pipeline. |
| FR-02 | Quality assessment | Score focus/sharpness, illumination, FOV coverage and artifacts; classify image as good, borderline or ungradable. |
| FR-03 | Recapture feedback | Reject ungradable images and provide a reason code plus actionable recapture instruction. |
| FR-04 | Adaptive enhancement | Apply CLAHE, illumination normalization and denoising only to borderline images; re-check quality afterward. |
| FR-05 | Retinal analysis | Provide vessel segmentation and, where feasible, optic-disc/fovea localization and lesion analysis. |
| FR-06 | DR grading | Predict ICDR-style severity Grade 0–4 and derive the competition-defined referable DR decision (Level 2+). |
| FR-07 | Explainability | Display Grad-CAM plus available lesion-level evidence and locations. |
| FR-08 | Confidence | Provide calibrated confidence and an uncertainty indicator; avoid presenting raw softmax as clinical certainty. |
| FR-09 | Human review | Allow ophthalmologist approval/override and preserve the final human decision. |
| FR-10 | Report | Generate a concise annotated screening report suitable for rapid review. |
| FR-11 | Simulation | Model acquisition rate, transmission/bandwidth, AI processing, queues and reviewer capacity in Simulink. |
| FR-12 | Validation | Produce reproducible metrics, confusion matrices, ROC/AUC, sensitivity/specificity and ablation results. |

# 6. AI / Technical Stack

| Technology | Use |
| --- | --- |
| Image Processing Toolbox | CLAHE (adapthisteq), denoising, illumination normalization, morphology |
| Deep Learning Toolbox | Transfer-learning DR classifier; U-Net/segmentation where justified; Grad-CAM |
| Computer Vision Toolbox | Geometric/feature methods and optic-disc/fovea localization fallbacks |
| Statistics and Machine Learning Toolbox | Calibration, ROC/AUC, sensitivity/specificity and statistical evaluation |
| Medical Imaging Toolbox | Medical image workflow/visualization as applicable |
| Simulink + SimEvents | Discrete-event modelling of arrival, queues, processing, network and review capacity |
| MATLAB UI / App Designer | Operator/ophthalmologist demo interface and report workflow |

# 7. Dataset & Experimental Strategy

APTOS 2019: primary DR grading training/validation/testing.

IDRiD: lesion-level analysis and additional grading validation.

DRIVE: vessel segmentation.

Messidor-2: untouched external validation set; do not use for training or hyperparameter tuning.

Use patient/eye-level separation where metadata permits; avoid crop-level leakage.

Use realistic augmentation and class-weighted/focal loss for class imbalance.

Report both multiclass grading performance and the Level 2+ referable target.

Do not claim >90% sensitivity / >85% specificity until experimentally demonstrated.

# 8. Explainability & Safety Requirements

Grad-CAM must be presented as model attention, not proof of causality.

Lesion evidence should be separated from classifier output unless validated as a predictive feature.

Low-confidence/uncertain cases must be routed for human review.

The prototype must clearly state that it is a screening/decision-support system, not a replacement for an ophthalmologist.

Sub-pixel microaneurysm detection is a target from the problem statement; it must not be claimed as achieved without evidence.

# 9. Simulink Requirements

Inputs: patient arrival rate, image acquisition time, image size, bandwidth, transmission delay, AI processing time, recapture rate, referral rate, review time and number of reviewers.

Outputs: throughput, waiting time, queue length, reviewer utilization, bandwidth requirement, bottleneck and annual capacity.

Run what-if scenarios for different patient loads, connectivity and ophthalmologist counts.

Demonstrate whether the configured system can support the 100,000+ patients/year scenario.

# 10. MVP / Scope Priority

MUST: quality gate + recapture; DR classifier; Grad-CAM; calibrated confidence; report/UI; basic Simulink model; rigorous validation.

SHOULD: lesion detection/evidence; vessel segmentation; optic-disc/fovea localization.

IF TIME: advanced lesion segmentation, neovascularization, advanced uncertainty analysis.

Principle: protect the working core. Optional segmentation must not block grading, explainability, validation or the Simulink demo.

# 11. Key Differentiators

Quality-gated screening with immediate recapture feedback.

Clinical-evidence-oriented explainability combining model attention with lesion evidence.

Calibrated confidence and uncertainty integrated into human review.

Operational Simulink model for bandwidth, queue and specialist-resource optimization.

These are integration/workflow differentiators; standard CNN, U-Net, CLAHE and Grad-CAM should not be marketed as inventions.

# 12. Success Criteria

Working end-to-end MATLAB prototype from fundus image to report.

Target the PS requirement of >90% sensitivity and >85% specificity for referable DR, with honest measured results.

External validation on a held-out dataset.

Explainability output that can be inspected quickly by a reviewer.

Simulink model demonstrating throughput, bottlenecks and resource scenarios for 100,000+ annual patients.

Reproducible experiments and ablation showing whether added modules improve the system.

# 13. Key Risks & Fallbacks

Lesion segmentation underperforms → use high-recall candidate detection / evidence-only mode rather than blocking DR grading.

External performance below target → report honestly; tune only on training/validation data and prioritize clinically appropriate sensitivity.

Time constraints → prioritize quality + grading + explainability + calibration + report + Simulink.

Heavy model too slow → benchmark smaller backbones and select the best accuracy/latency trade-off.

# 14. Final Demo Flow

1) Upload/capture fundus image → 2) quality score → 3) show recapture for a bad image → 4) process a good image → 5) show DR grade → 6) show Grad-CAM + lesion evidence + confidence → 7) generate report → 8) ophthalmologist approve/override → 9) show Simulink district-scale scenario and bottleneck/resource outputs.

# 15. Product Principle

Detect → Explain → Validate → Refer → Scale. The product succeeds when every major component maps directly to the SIH requirements and can be demonstrated with measured evidence.